Flow Rate-Rise Time Profiles from Model Dry Powder Inhaler (DPI) Testing of Abbreviated Impactor Systems Compared with Their Full Resolution Counterparts: Initial Experimental Data

Mitchell JP1, Roberts DL2, Versteeg H3, Cooper A4, Copley M5 & Greguletz R6

1 Jolyon Mitchell Inhaler Consulting Services Inc., 1154 St. Anthony Road, London, N6H 2R1, Canada

2 Applied Particle Principles LLC, 17347 Westham Estates Court, Hamilton, VA 20158, USA

3 Mechanical and Manufacturing Engineering, Loughborough University, Loughborough, LE11 3TU, UK.

4 Kindeva Drug Delivery, Derby Road, Loughborough, LE11 5SF, UK

5 Copley Scientific Ltd., Colwick Quays Business Park, Rd. No 2, Colwick, Nottingham NG4 2JY, UK

6 Sofotec GmbH, a member of the AstraZeneca Group, Benzstraße 1-3, Bad Homburg, D-61352, Germany

Summary

We report the outcome of a European Pharmaceutical Aerosol Group study evaluating flow rate-rise times for reduced Next Generation Impactor (rNGI), Fast-Screening Impactor (FSI) and Fast-Screening Andersen impactor (FSA) with surrogate low-, medium- and high-resistance DPIs. Time for the flow rate to reach 90% of target (t90) and flow acceleration rate (slopet20/t80) defined the profiles. Volumes of the abbreviated impactors were estimated (±10mL) and compared to published values for their counterpart parent impactors (rNGI and FSI to NGI; FSA to Andersen Cascade Impactor (ACI)). t90 (n=3 replicates; mean±S.D.) for rNGI (system volume 1990 mL, participant D) for target flow rates of 30 and 60 L/min were 387±5 and 218±8 ms respectively, close to corresponding values of 400±7 and 219±6 ms for the NGI (2000 mL). The slopet20/t80) values were 0.09±0.00 and 0.31±0.00 L/min/ms at 30 and 60 L/min for the rNGI, also similar to the NGI value. Values of t90 (170±5; 76±2 ms) and slopet20/t80 (0.22±0.01; 0.96±0.01 L/min/ms) with the smaller FSI (1150 mL), were significantly different at 30 and 60 L/min respectively than those of the NGI.  Data from Participant G contained similar divergences including tests at 90 L/min (t90 60±3 ms (FSI); 185±6 ms (NGI).  When testing two FSAs (volume of 630 mL) at 60 L/min, Participant J found almost identical t90 values (70±3; 71±2 ms), shorter than 60-L/min values (151±3; 162±6 ms) for their two ACIs (larger volume, 1180 ml). Apparatus internal volume is an important determinant of flow rate-rise time.

Key Message

Our study has identified that the FSI and FSA abbreviated impactors start up more quickly than their parent apparatuses whereas the rNGI starts up nearly at the same rate as the NGI.  Apparatus internal volume is an important determinant of the flow rate-rise time profile.

Introduction

Flow rate-rise time profiles obtained when testing DPIs following procedures in the pharmacopeial compendia are important in vitro characteristics of likely clinical performance because they affect both powder-to-aerosol generation and transport processes governing medication delivery efficiency [,]. An EPAG cross-industry study investigated these profiles with full resolution cascade impactors at target flow rates of 30, 60 and 90 L/min and with orifice plates representative of high-, medium-, and low-resistance DPIs []. Three of the nine original participating groups (blinded as the letters D, G and J) also tested abbreviated impactors.  These data offered the opportunity to explore the impact of differing internal volume between reduced and parent full-resolution impactor, an issue identified as a potential source of discrepancies in measures of fine particle mass between at least one abbreviated and full resolution impactor system []. The intrinsic resistance to flow of the apparatus, arising mostly from the lower stages of the impactor and back-up filter/micro-orifice collector, also contributes to the flow rate-rise time behaviour, but was not explored explicitly. The goal of this study was to quantify how such differences may affect the characteristic flow rate-rise time profiles by comparing flow rate start-up time and acceleration rate immediately after the flow initiation by the opening of a solenoid valve downstream of the apparatus as is specified in pharmacopeial procedures [].

Materials and Methods

The experimental methodology followed closely the pharmacopeial procedure for APSD determination from DPIs, as reported previously [3] for the flow rate-rise time study involving the full resolution impactors. This procedure was also undertaken by participants D, G and J, who provided measurements with reduced NGI (rNGI) [], Fast Screening Impactor (FSI) [] and Fast Screening Andersen impactor (FSA) []. The apparatuses that were evaluated together with their nominal internal volumes are summarized in Table 1.  Measurements (n=3 replicates) were made of the flow rate-rise time profiles at target flow rates of 30, 60 and 90 L/min with a flow meter that was immediately downstream of a purpose-built orifice-plate surrogate DPI [3] designed to cause the flow rate to reach the chosen target when its pressure drop reached 4kPa.

Table 1: Apparatus Configurations and Nominal Internal Volumes (all ±10 mL)

| Participant | Full Resolution |  |  | Abbreviated |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Impactor | Pre-separator present | Total internal volume (mL) | Impactor | Pre-separator present | Total internal volume (mL) |
| D* | NGI | Y | 2000 | rNGI FSI | Y N | 1990 1150 |
| G** | NGI | Y | 1990 | FSI | Y | 1140 |
| J | ACI*** (2) | Y | 1180 | FSA (2) | Y | 630 |

* Additional volume for interconnections was assumed to be 40 mL

** Additional volume for interconnections was specified as 25 mL

*** 28.3 L/min configuration

Surrogate DPIs were constructed with aperture dimensions causing a 4-kPa pressure drop that simulated either a high-, medium-, or low-resistance DPI at the target flow rates of 30, 60, and 90 l/min [3]. Total internal volumes for the full resolution impactors were as provided in Greguletz et al.3, rounded to three significant digits, and included allowances for the pre-separator, USP/PhEur induction port (67 mL), the flow meter (45 mL), the volume downstream of the orifice plate assembly (10 mL) and, unless specified, 40 mL for the interconnecting tubing upstream of the flow controller solenoid valve. Upon actuation of this solenoid valve to initiate the sampling and measurement processes, each flow rate-rise time profile was recorded at 5 ms intervals from the signal provided by the flow meter (model 4040, TSI Inc., Shoreview, MN, USA) by purpose-developed proprietary recording software. The resulting data enabled the calculation of three characteristic rise time indicators, t20, t80, and t90 (each in ms), representing times to attain 20%, 80%, and 90% of the steady-state target flow rate.  t90 and the slope of the profile (slopet20/t80), calculated by linear interpolation between t20 and t80, were therefore the primary measures.

Results

The t90 values for both abbreviated and full-resolution parent impactors are presented in terms of

Figure 1 -	t90 Values for Abbreviated and Full-Resolution Impactors Derived from Flow Rate-Rise Time Profiles with Target Flow Rates Mimicking Low-, Medium- and High-Resistance DPIs

apparatus internal volume and target flow rate in Figure 1. Thus, participant D, who submitted data for both the rNGI and FSI have results linked with their NGI. Importantly, t90 values for the rNGI at both target flow rates evaluated by this participant (30 and 60 L/min) were very close to their reference value for the NGI and were associated with comparable internal volumes close to 2000 mL. In contrast, t90 values for the FSI and FSA systems were significantly shorter than the corresponding values obtained with their corresponding full-resolution NGI and ACI configurations respectively.

The slopet20/t80 values are similarly presented for each configuration in Figure 2. The smallest values of this metric were associated with the slowest flow rate-rise profiles, and as expected, occurred with the apparatuses having the largest internal volumes. Again, the values for the rNGI were very close to those for the NGI (Participant D). Those for the FSI apparatuses evaluated by participants D and G were in close agreement with each other where direct comparison was possible (target flow rates of 30 and 60 L/min. However, they were significantly larger than the corresponding values for their NGIs).  Values for slopet20/t80 for the two FSA apparatuses evaluated by participant J at each target flow rate were also very similar to each other, but significantly greater than equivalent values for the two parent ACIs.

Figure 2 – Slopet20/t80 Values for Abbreviated and Full-Resolution Impactors Derived from Flow Rate-Rise Time Profiles with Target Flow Rates Mimicking Low-, Medium- and High-Resistance DPIs

Discussion

Comparisons of both t90 and slopet20/t80 values for abbreviated impactor systems and their full resolution parent impactor counterparts highlighted the importance of apparatus internal volume in governing the shape of the flow rate-rise time profiles.  An associated abstract [] in which a validated physical model of the underlying processes taking place during the time when the instantaneous flow rate is unsteady when evaluating DPIs [] provides support that apparatus internal volume governs the rate at which target flow is achieved.  Changes in both t90 and slopet20/t80 were directly associated with differences in the magnitude of the “trapped volume” (equivalent to the system total internal volumes in the present study) behind the surrogate DPI. As an example, relating to the APSD measurement outcome, the systematic over-estimate of fine particle mass fraction (<5 μm aerodynamic diameter) reported by Russell-Graham et al., comparing FSI with parent NGI measurements in association with DPI testing [] may be explicable by the availability of more energy for powder dispersion associated with the steeper flow rate-rise time curve shortly after opening the solenoid valve to initiate flow into the apparatus.

It should be appreciated that the results reported are a by-product of the much larger flow rate-rise time study [3] and are based on a very limited number of replicate measurements at each condition. They also do not address specifically the role that apparatus resistance may play in determining the flow rate-rise time profile. A designed experiment powered to enable statistical analysis of the comparisons between abbreviated and parent full resolution impactors should be undertaken to validate them and to address explicitly the contribution made by apparatus resistance.  Nevertheless, our study has provided support for a proposal that in order to achieve improved agreement between abbreviated and full-resolution parent impactor, either use systems with similar internal volume, such as the rNGI/NGI pairing, or add sufficient internal volume to the abbreviated system to match that of the parent impactor.

Finally, it should be emphasized that these findings apply exclusively to the testing of DPIs by the procedures in the pharmacopeial compendia, where flow is initiated at the start of measurement, rather than having a constant flow rate through the apparatus, as intended for ideal size-fractionating performance [].

Conclusions

We conclude that apparatus internal volume is a key determinant of flow rate-rise time profiles when evaluating DPIs in accordance with the methods for APSD measurement in the pharmacopeial compendia.

Acknowledgements

We acknowledge the participants D, G and J for their data that has made these comparisons between abbreviated and full-resolution impactor systems possible.

References

[] Chavan V, Dalby R. Effect of rise in simulate inspiratory flow rate and carrier particle size on powder emptying from dry powder

inhalers. AAPS PharmSciTech 2000; 2 (2): pp. 1–7.

[] Coates M S, Chan H-K, Fletcher DF, Raper JA. Influence of air flow on the performance of a dry powder inhaler using

computational and experimental analyses. Pharm Res 2005; 22 (9): pp. 1445–1453.

[] Greguletz R, Andersson PU, Cooper A, Chambers F, Copley MA, Daniels G, Hamilton M, Hammond M, Mohammed H, Roberts

DL, Shelton C, Versteeg HK, Mitchell JP A cross-industry assessment of the flow rate-time profiles of test equipment typically

used for dry-powder inhaler (DPI) testing: Part 1 –Compendial apparatuses. Aerosol Sci Technol. 2020; 54(12): pp.1424-1447.

[] Copley M, Mitchell JP, Svensson M, Christopher JD, Quiroz J, Daniels G, Hamilton M, Russell Graham D. Validating AIM-

Based Instrumentation and Associated Measurement Techniques. In: TP Tougas, JP Mitchell, SL Lyapustina (eds): Good

Cascade Impactor Practices, AIM and EDA for Orally Inhaled Products, Springer, New York, NY, USA; pp. 283-358. 2013.

[] European Directorate for Quality in Medicines and Healthcare (EDQM). 2019. European pharmacopoeia 10.0,

Monograph 2.9.18. Preparations for inhalation: Aerodynamic assessment of fine particles. EDQM, Strasbourg, France.

[]  Svensson M, Berg E. Measuring fine particle dose using inter-stage filters in the NGI: An overview of the methods. (Abstract). Presented at: Drug Delivery to the Lungs. The Aerosol Society, Edinburgh, UK. 2010; 21: pp. 382-385.

[] Roberts D L, Romay FJ, Design of the Fast-Screening Impactor Based on the NGI Pre-Separator.  (Abstract). Presented at: Drug Delivery to the Lungs, The Aerosol Society, Edinburgh, UK. 2009; 20: pp.206-209.

[] Mitchell JP, Nagel MW, Doyle C, Ali RS, Avvakoumova V, Christopher D, Quiroz J, Strickland H, Tougas T, Lyapustina S.

Relative precision of inhaler aerodynamic particle size distribution (APSD) metrics by full resolution and abbreviated Andersen

Cascade Impactors (ACIs): Part 1. AAPS PharmSciTech 2010; 11(2): pp. 843–851.

[] Versteeg HK, Roberts DL, Cooper A, Mitchell JP. A model of the rise time and flow acceleration of AIM impactors for the testing

of Dry-Powder Inhalers. (Abstract). Presented at: Drug Delivery to the Lungs 2022, The Aerosol Society, Edinburgh UK. 2022; 33: not yet assigned.

[] Versteeg HK, Roberts DL, Cooper A, Chambers F, Copley M, Mitchell. JP A cross-industry assessment of the flow rate-

elapsed time profiles of test equipment typically used for dry-powder inhaler (DPI) testing: Part 2. –Analysis of transient air

flow in the testing of DPIs with compendial cascade impactors. Aerosol Sci. Technol. 2020; 54(12): pp. 1448-1470.

[] Russell-Graham D, Cooper A, Stobbs B, McAulay E, Bogard H, Heath V, Monsallier E. Further evaluation of the fast screening

impactor for determining fine particle fraction of dry powder inhalers. (Abstract). Presented at: Drug Delivery to the Lungs-21, The Aerosol Society, Edinburgh, UK. 2010; 21: pp. 374–377.

[] Marple VA, Liu BYH. Characteristics of laminar jet impactors. Environ Sci Technol. 1974; 8: pp. 648–654.